| The Diabetes Care in the Community for Support Workers course teaches support workers how to support people with a disability and diabetes. Module 2 of the course focuses on blood glucose monitoring. A support worker needs to be able to accurately and safely check the blood glucose level of a person with diabetes when in their care. To learn how to use the meter, click on the link and watch the training video for the meter you use. Read through the procedure below and use the meter and quality control solution provided to practice the recommended technique. This helps you to achieve an accurate technique, while also preparing you for the evaluation component of this Module. Jayne Lehmann RN CDE will check your blood glucose monitoring procedure during an Internet video-linked session. Your organisation will tell you the date, time and venue for the session. |  |  |  |
| --- | --- | --- | --- |
| Procedure: How to use a blood glucose meter |  |  |  |
| Type of meter: [ ] Accu-Chek Guide [ ] Accu-Chek Performa |  |  |  |
| Name of your organisation: |  | Your name: |  |
| Name of the service in which you work: |  | Today’s date: |  |
| Step | Description |  | Done |
| - Gather equipment | Meter, bottle of strips, finger pricker safe for support worker to use, gloves, cotton wool ball, quality control solution, sharps container, chart/logbook/BGM order, pen, damp flannel or hand washing facilities, sharps disposal unit. |  |  |
| - Prepare equipment | Wash your hands. Put gloves on. Wash/wipe the client’s hands. Put strip in meter – close strip bottle lid as soon as strip is removed. Get the finger pricking device is ready to use. |  |  |
| - Do blood glucose check* | Use a safe finger-pricking device to prick the side of the finger. Squeeze finger from the base, moving the blood down the finger. Bring the meter and strip towards the finger so that the end of the strip touches the drop of blood. The meter starts to count down when sufficient blood is applied to the strip. Give the person a cotton wool ball and ask them to push this over the finger-prick site for 1 minute. The number on the screen = blood glucose level. |  |  |
| *Use quality control solution NOT BLOOD to practice your technique and for audit. | Squeeze a drop of quality control solution onto the side of your finger. Bring the meter towards the finger and touch the end of the strip onto the drop of blood. Wipe excess quality control solution from finger. The number on the screen = quality control result. |  |  |
| - Interpret result | Blood glucose level: What does the number mean – low, healthy or high? Use the Check, Think and Act resource to help decide. Quality control solution: Check number on side of the bottle of strips to see if it falls within the acceptable range for the solution. |  |  |
| Blood glucose meter problem solving and maintenance |  |  |  |
| Step | Description |  | Done |
| - Codes | If a code appears on the meter how do you respond to it? You look in the meter’s manual to identify the meaning of the code and respond as directed to rectify the problem. |  |  |
| - Meter maintenance | How is the meter maintained? a) Batteries changed b) Blood spills wiped up immediately c) If meter is dropped do a quality control check. |  |  |
| 7. Record BGL | Write the result on the blood glucose monitoring chart/log book or medication chart depending on your organisation’s procedure. |  |  |
| 8. Problem solving Can call the Accu-Chek Enquiry Line: 1800 251 816 (toll free) for assistance too. |  |  |  |
| a) Meter stops working | Check the batteries are in correctly. Still not working – put in fresh batteries. Still not working - call company FREECALL number for further assistance. |  |  |
| b) Unexpected result | Wash the person’s hands. Repeat blood glucose check and/or use quality control solution to check meter’s accuracy by putting a drop of solution into strip instead of blood. Check result against the expected range for that control solution. This is found on the side of the strip bottle. Result in-range = accurate. Result out of range = call Company FREECALL number. Accu-Chek Enquiry Line: 1800 251 816 (toll free) |  |  |
| c) Quality control | Use the quality control solution to check the meter’s accuracy by using it instead of blood. Check once a week. Change the person checking the result each week so that everyone who checks blood glucose levels using this meter get to do the quality control checks from time-to-time. After opened, the solution lasts 3 months. |  |  |
| d) Pricked by a used needle. | Immediately wash the site thoroughly with soap and running water. Report incident to Manager. Go and see the doctor for a medical review. N.B. Follow your organisation’s specific procedure. |  |  |
| Remember to successfully pass Module 2: Blood Glucose Monitoring you need to achieve 80% + in the on-line test and pass the check of your blood glucose monitoring technique by Jayne Lehmann. Use of the finger-pricking device for blood glucose monitoring It is not appropriate for support workers to use a finger-pricking device that requires you to remove the used and exposed lancet with your fingers. This substantially increases the risk of a needle-stick injury. Use a finger-pricking device that never has the needle left exposed after use. Suitable options include: - Single use, disposable finger pricking device - Roche Fast Clix or Multi-Clix finger pricking device with replaceable lancet cartridges (available from the local pharmacy). |  |  |  |
| - Preparation | Wash your hands and put on gloves. Prepare equipment, load the finger-pricker as per instructions from the video. Wash the person’s hands. Explain how you will prick their finger on the side – it doesn’t hurt as much - using the finger pricker. |  |  |
| - Finger prick site | Ask the person which finger they would like you to prick and do it on the side so it hurts less. Squeeze the finger from the base, moving the blood to the tip of the finger. |  |  |
| - Blood application | Pick-up the meter with a strip already inserted and check it is ready for the crop of blood. Touch the end of the strip against the blood drop on the finger. Ask the person to push a cotton wool ball firmly over the finger prick site for 1 minute to prevent bruising. |  |  |
| 4. Used lancet disposal | Single-use device: Throw the whole device into the yellow sharps disposal unit once used. Fast-Clix of Multi-Clix finger pricking device: There are 6 lancets in the cartridge. After pricking the finger, use the lever to turn the cartridge to a new lancet. The number on the side of the device will, for example, change from 6 to 5 when moving the lever for the first time with a new lancet cartridge. Once it shows 0 on the side of the device, take the end cap off and remove the used cartridge with your fingers. This is safe because all of the used lancets are covered. Throw this into the yellow sharps disposal unit. |  |  |
| 5. Tidy up | Record result, wash hands and put everything away. |  |  |
| Please note you will not be evaluated in the use of the finger-pricking device when Jayne Lehmann checks your technique with the meter. This information is provided for you to read and self-evaluate. |  |  |  |